Table 1. Detection sensitivity across five CT acquisition conditions (21 cases, 126 consensus nodules, ≥3 readers, confidence threshold = 0.5, match distance = 15 mm).

| Imaging Condition | Detected | Total Nodules | Sensitivity (%) |
| --- | --- | --- | --- |
| Baseline | 57 | 126 | 45.2% |
| 25% Dose | 52 | 126 | 41.3% |
| 50% Dose | 53 | 126 | 42.1% |
| 3mm Thick | 52 | 126 | 41.3% |
| 5mm Thick | 33 | 126 | 26.2% |

pp = percentage points. The 5mm slice thickness condition (highlighted) produced the largest sensitivity reduction: −19 pp from baseline, representing a 42% relative decrease.